## Slide 1
Bioengineering Senior Design | Spring 2017
Aras Fanuscu, Andrea Frank, David Hernandez, Angel Xiao

## Slide 3: The Numbers
23,000
people in the US sprain their ankles every day
of people don’t complete the rehabilitation process
Chart: 
| 1st Qtr | 2nd Qtr
Column1 | 65 | 35
65%
of people experience recurrence of their injury
Chart: 
| 1st Qtr | 2nd Qtr
Column1 | 38 | 62
38%

## Slide 4: The Rehabilitation Process
Initial Treatment
Rehabilitation
Functional Strengthening
The SockRocker Rehabilitation Device

## Slide 5: The Rehabilitation Process
Initial Treatment
Inversion/Eversion
The SockRocker Rehabilitation Device

## Slide 6: Bridging the Gap

## Slide 8: We set three main Goals
Effective rehabilitation
Targeted muscle therapy
Active rehabilitation
Personalization
Physician input
Variable resistance
Accessibility
Universal design
Useful in many contexts

## Slide 9: The Design Process
The SockRocker
Design Drivers
Physiological Constraints
Performance Requirements
Experienced Professional Input
Range of Motion
Variable Resistance
Physician Control
Comfort

## Slide 11
SR

## Slide 12: Physician-patient Information Flow
Physician Input (Matlab GUI)
Arduino Microcontroller sends Signal
Varying current input to brake
Brake generates resistive force
Patient progress tracked and recorded
Patient carries out assigned training

## Slide 13: Physician-patient Information Flow
Physician Input (Matlab GUI)
Arduino Microcontroller sends Signal
Varying current input to brake
Brake generates resistive force
Patient progress tracked and recorded
Patient carries out assigned training

## Slide 14: Variable resistance system
Varying current input to brake
Varying current modulates the brake’s resistance to rotation
Physician Input (Matlab GUI)

## Slide 15: Physician-patient Information Flow
Physician Input (Matlab GUI)
Arduino Microcontroller Sends PWM Signal
Varying current input to brake
Brake generates resistive force
Patient progress tracked and recorded
Patient carries out assigned training

## Slide 16: Our track ensures Targeted Therapy
Isolates inversion and eversion motions
Shape determined by ankle sizes and motion
Patient carries out assigned training

## Slide 18: Effective rehabilitation
0.8— 5.3 N-m
Torque range
0 — 20 N-m
Torque resolution
0.5 N-m
Range of Motion
-23o — 12o
-30o — 30o
0.09 N-m

## Slide 19: Personalized rehabilitation
Physician prescribed
force profiles
Patient customization
Monitor progress
Save & Send

## Slide 20: Increasing Accessibility to care
Ease of
Acquisition
Ease of
Usage
Low Cost
$360 per device
Comfort
4.8/5 rating (pilot study)
Portable
<15lb, no external wires
Low Power
24V battery: 1 month
Universal Design
Fits any foot size
Human factors
Familiar interface

## Slide 22: We see a bright future for our device
Improve range of feet sizes and forces
Incorporate more sophisticated controls
Clinical testing to affirm long term efficacy
Develop activities to improve compliance

## Slide 23: Our Impact
7.2 M
injuries
900,000
ER visits
$1000 per visit
$900 MM / yr
$360/device

## Slide 24: Our Impact
“ Millions of Americans each year are incompletely rehabilitated and prone to re-injury, due to lack of effective exercise equipment to specifically strengthen key ankle stabilizer muscles.
This prototype device illustrates a way to address this unfortunate gap in our clinical rehabilitation management of [these] persons.”
Dr. Timothy Dillingham, MD, MS
Chair, Dept. of Physical Medicine and Rehabilitation

## Slide 25
Thank you to the BE Department for all their support!
Special Thanks to:
Dr. Timothy Dillingham
Dr. Michael Rizk
Dr. Susan Marguiles
Sevile Mannickarottu
Daniel Pak

## Slide 27
Assembly Drawing
Overview Images
Track Details
Force Interfaces
Circuit Diagram
Force Calibration
Detailed Costs
References
Appendix

## Slide 28: Assembly Drawing

## Slide 29: Overview Images

## Slide 30: Track Details

## Slide 31: Force Interfaces

## Slide 32: Circuit Diagram
Not shown: Bluetooth Receiver Section

## Slide 33
Chart: Circut Resonse Curve
| 0.01 | 10.01 | 20.009999999999998 | 30.009999999999998 | 40.01 | 50.01 | 60.01 | 70.009999999999991 | 80.009999999999991 | 90.009999999999991 | 100.00999999999999
Average Voltage Output | 0.64367 | 2.911 | 4.026 | 4.9023 | 5.6273 | 6.1887 | 6.7937 | 7.3677 | 7.8568 | 8.305 | 8.749
Predicted Voltage Output | 0 | 2.3326 | 4.0607 | 5.3409 | 6.2893 | 6.9918 | 7.5123 | 7.8979 | 8.1835 | 8.3952 | 8.5519
Chart: Average Force Requirements
| 0 | 25 | 50 | 75 | 100
Avg | 9.8797 | 13.8607 | 14.8543 | 13.8607 | 19.06

## Slide 34: Cost Breakdown
Electronics | Qty | Cost per Unit | Total Cost
Arduino | 2 | $17.99 | $35.98
Bluetooth Set | 2 | $4 | $8
Brake | 1 | $130.35 | $130.35
9V Batteries | 2 | $4.5 | $9
Misc. Electronics | 1 | $1 | $1
SUM: |  |  | $184.33
Mechanical | Qty | Cost per Unit | Total Cost
Pulleys | 4 | $7.02 | $28.08
Wire Rope | 3 | $8.39 | $25.17
Ball Bearings | 5 | $6.03 | $30.15
Clamps | 2 | $5.28 | $10.56
Foot Piece | 1 | $20.50 | $20.50
Leg Piece | 1 | $30.72 | $30.72
3D Printing Costs | 586.02 | $0.05 | $31.25
Laser Cut Costs | 4 | $18.80 | $75.20
SUM |  |  | $176.4344
Grand Total: 360.7644

## Slide 35
Konradsen, L., Olesen, S., & Hansen, H. M. (1998). Ankle sensorimotor control and eversion strength after acute ankle inversion injuries. The American journal of sports medicine, 26(1), 72-77.
Osteopathicinfo. Melbourne Osteopathy Group - Osteopath Fitzroy North. 23 Jan 2017. Retrieved from: <http://www.osteoinfo.com.au/self-management/management-of-the-ankle-foot>.
Bernier, J. N., Perrin, D. H., & Rijke, A. (1997). Effect of unilateral functional instability of the ankle on postural sway and inversion and eversion strength. Journal of Athletic Training, 32(3), 226.
Goncalves, F. D. (2005). Characterizing the behavior of magnetorheological fluids at high velocities and high shear rates (Doctoral dissertation, Virginia Tech).
Lord Rheonetic Magnetically Responsive Technology, Hydrocarbon-Based MR Fluid MRF-132LD, Product Bulletin, 2002.
http://cruxcrush.com/2014/11/11/climber-problems-ankle-injury/
http://www.laboratoryequipment.com/news/2013/10/anklebot-can-aid-stroke-rehab
http://medicalrecruiting.com/wp-content/uploads/2014/02/Physician-at-computer.jpg
http://images.wisegeek.com/circuit-board.jpg
http://www.electric-brake.com/img/magnetic-partice-brake.png
http://skpang.co.uk/catalog/images/sensors/accelerometer/10121-01.jpg
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC3183932/
References